## Design Input Requirements
| Lewis Lab Origami Face Shield DESIGN INPUT REQUIREMENTS | [5 empty] | REV:A
Page 1/1
Design Input ID | Design Input Description/Requirement | Rationale | User Need Trace | Potential Test Method | Source | Notes
DI-FS-01 | Protect eyes, nose, mouth from splashes, sprays, spatter, and body fluids | The face shield must protect the user's eyes, nose, and mouth | Protect user's face from splashes and sprays | Face Shield Spray Test | ISO 18526-3:2020
DI-FS-02 | Avoid potential splash to drip down from forehead | The face shield must protect the user's eyes, nose, and mouth | Protect user's face from splashes and sprays | Face Shield Spray Test | ISO 18526-3:2020
DI-FS-03 | Optical quality (non-glare, non-fogging, anti-static properties) visor that does not impede vision, peripheral vision | The user should be able to see through the shield and be able to perform normal functions | No interference with user's vision | Face Shield Optical Quality Test | CASMS Vision Acuity Test Protocol
DI-FS-04 | Biocompatible materials not causing user skin irritation | The face shield should not irritate the user in any way | User can comfortably use for a 8-12 hour shift | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
DI-FS-05 | Compatible with stethoscope | No itteruption of normal functions | Does not interfere with user's normal functions | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
DI-FS-06 | The User must be able to don and doff the shield easily | The user should be able to easily don and doff the face shield to prevent potential exposure to virus |  | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
DI-FS-07 | Compatible with use of glasses or goggles | No itteruption of normal functions | Does not interfere with user's normal functions | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
DI-FS-08 | Compatible with off the shelf disposable surgical mask or respirator use | No itteruption of normal functions | Does not interfere with other PPE | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
DI-FS-09 | User must be able to perform normal movements with shield on | No itteruption of normal functions | Does not limit user's range of motion | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
DI-FS-11 | Face Shield must be able to be worn comfortably for an 8-12 hour shift | User's may need to wear the face shields for an entire shift of 8-12 hours | User can comfortably use for an 8-12 hour shift | Face Shields Human Factors Test | FDA-Applying Human Factors and Usability Engineering to Optimize Medical Device Design, Feb 3, 2016
Labeling
DI-FS-12 | The product is labeled accurately to describe the product as a face shield for medical purposes | FDA EUA guidance | Face Shield must be properly labeled | Face Shields Package Insert Requirement | FDA EUA guidance 4/13/2020
DI-FS-13 | Lists body contacting materials | FDA EUA guidance | Face Shield must be properly labeled | Face Shields Package Insert Requirement | FDA EUA guidance 4/13/2020
DI-FS-14 | Includes labeling that describes the product as intended for either a single-user, single use, or for multiple uses by the same user | FDA EUA guidance | Face Shield must be properly labeled | Face Shields Package Insert Requirement | FDA EUA guidance 4/13/2020
DI-FS-15 | Includes instructions for recommended cleaning and/or disindection materials and processes if applicable | FDA EUA guidance | Face Shield must be properly labeled | Face Shields Package Insert Requirement | FDA EUA guidance 4/13/2020
DI-FS-16 | The face shield does not contain any materials that will cause flammability, or the product meets Class I or Class II flammability requirement per 16 CFR 1610 (unless labeled with a recommendation against use in the presence of high intensity heat source or flammable gas) | FDA EUA guidance | Face Shield must be properly labeled | Face Shields Package Insert Requirement | FDA EUA guidance 4/13/2020
DI-FS-17 | The product is not intended for any use that would create an undue risk in light of the public health emergency; for example, the labeling does not state that use of the authorized face shield alone will prevent infection from microbes or viruses, or that it is effective against radiation protection. | FDA EUA guidance | Face Shield must be properly labeled | Face Shields Package Insert Requirement | FDA EUA guidance 4/13/2020

## READ ME
| DESIGN INPUT REQUIREMENTS |  |  | REV:D
2020-04-29
Note: This document is intended to provide a template for developing and documenting the design inputs associated with your device design. This SHOULD be prototype or design specific, with the intention of each desgin being to fulfill part of or all of a specific user need. DISCLAIMER:This content is made available subject to the following conditions: This information (the “Information”) is made available for educational and research purposes only. It has not been reviewed or approved by the United States Food and Drug Administration or any other agency. The Information is provided “AS IS” and must not be used to make a clinical diagnosis, to provide clinical care, or to replace or overrule a licensed health care professional’s judgment. Mass General Brigham, including any and all of its affiliated institutions, is not responsible for ensuring that one’s use of the Information will be clinically sound, without error, or otherwise successful. If you choose to use the Information, you assume all risk and bear all responsibility and liability. Mass General Brigham specifically disclaims all warranties, express, implied or statutory, including implied warranties of merchantability, fitness for a particular purpose, and noninfringement.
Author: | Evan Sevieri | Reviewed By: (MGB Central) | [Reviewer Name and Date]